LIST OF SUPPLEMENTARY MATERIALS

- SPECIES IDENTIFICATION

| Amplification of the rbcL gene M: 100 bp DNA Ladder 1: Sample A 2: Sample B (+): Positive control samples | # DNA sequencing # ACAAATTGAATTATTATACTCCTGACTACGAAGTCAAAGATACTGATATCTTGGCAGCATTCCGAGTAACTCCTCAACCTGGAGTTCCACCCGAAGAAGCAGGGGCTGCGGTAGCTGCCGAATCTTCTACTGGTACATGGACAACTGTGTGGACTGATGGACTTACCAGTCTTGATCGTTACAAAGGGCGATGCTACCACATAGAGGCTGTTATTGGGGAGGATAATCAATATATTGCTTATGTAGCTTATCCTTTAGACCTTTTTGAAGAAGGTTCTGTTACTAACATGTTTACTTCCATTGTGGGTAATGTGTTTGGTTTCAAAGCCTTACGAGCTCTACGTTTGGAGGATCTGCGAATTCCCACTTCCTATTCCAAAACTTTCCAAGGCCCGCCTCACGGCATTCAGGTTGAAAGAGATAAGTTGAACAAGTATGGTCGTCCCCTATTGGGATGTACTATTAAACCAAAATTGGGATTATCTGCAAAAAACTACGGTAGAGCAGTTTATGAGTGTCTAC |
| --- | --- |
| Comparing to the GenBank sequence database |  |
| Conclusion: The analyzed sample sequence was confirmed with 100% similarity with Curcuma longa species, compared to the GenBank sequence database |  |

- OPTIMIZATION OF LC CONDITION

B1. Stationary phase

(a) C8 column, (b) Phenomenex Synergi Hydro-RP column, (c) C18 column

B2. Mobile phase

| Methanol – 1% AF (50B:50D) | Methanol – H2O (50B:50C) |
| --- | --- |
| ACN : 1% AF (50A:50D) | ACN : H2O (50A:50C) |

B3. Buffer solution

| Buffer solution |  | 0.1 % AA | 0.5 % AA | 1 %AA | 0.1 % AF | 0.5 % AF | 1 % AF |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Retention time | CUR | 14.33 | 14.11 | 14.29 | 14.29 | 14.35 | 14.59 |
|  | DMC | 12.94 | 12.75 | 12.91 | 12.90 | 12.95 | 13.16 |
|  | BDMC | 11.69 | 11.53 | 11.66 | 11.65 | 11.69 | 11.88 |
| Area | CUR | 21,364,665 | 21,221,173 | 36,641,585 | 9,971,834 | 20,682,064 | 20,989,973 |
|  | DMC | 17,407,517 | 17,558,071 | 36,300,696 | 17,073,973 | 17,349,185 | 17,336,993 |
|  | BDMC | 10,146,657 | 10,381,599 | 19,433,646 | 9,971,834 | 10,107,171 | 10,248,430 |
| Asymmetry | CUR | 1.25 | 1.26 | 1.26 | 1.26 | 1.26 | 1.26 |
|  | DMC | 1.29 | 1.28 | 1.27 | 1.24 | 1.27 | 1.31 |
|  | BDMC | 1.21 | 1.21 | 1.21 | 1.19 | 1.21 | 1.23 |
| Resolution | CUR | 3.26 | 3.22 | 3.15 | 3.224 | 3.26 | 3.36 |
|  | DMC | 3.17 | 3.15 | 3.18 | 3.18 | 3.189 | 3.28 |
|  | BDMC | - | - | - | - | - | - |

B4. Mobile phase ratio

| Ratios |  | 40:60 | 50:50 | 60:40 |
| --- | --- | --- | --- | --- |
| Retention time | CUR | - | 14.29 | 4.75 |
|  | DMC | 37.16 | 12.91 | 4.45 |
|  | BDMC | 32.87 | 11.66 | 4.18 |
| Area | CUR | - | 36,641,585 | 33,965,917 |
|  | DMC | 9,005,631 | 36,300,696 | 32,087,362 |
|  | BDMC | 17,160,359 | 19,433,646 | 17,748,731 |
| Asymmetry | CUR | - | 1.26 | 1.32 |
|  | DMC | 1.21 | 1.27 | 1.29 |
|  | BDMC | 1.06 | 1.21 | 1.19 |
| Resolution | CUR | - | 3.15 | 1.80 |
|  | DMC | 3.87 | 3.18 | 1.64 |
|  | BDMC | - | - | - |

- OPTIMIZATION OF EXTRACTION PROCESS

C1. Solvent extraction

| Solvent extraction | CUR |  | DMC |  | BDMC |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Area | Content | Area | Content | Area | Content |
| Methanol | 17384059 | 4.09 | 5787434 | 1.36 | 6142869 | 1.44 |
| Ethanol | 14148233 | 3.33 | 4625665 | 1.09 | 4855062 | 1.14 |
| Aceton | 7528710 | 1.77 | 2367985 | 0.56 | 2540346 | 0.6 |
| Ethylacetate | Loss of resolution |  |  |  |  |  |

C2. Ratio of medicinal herb/solvent

| Ratio | CUR |  | DMC |  | BDMC |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Area | Content | Area | Content | Area | Content |
| 1:10 | 9386906 | 2.21 | 3066948 | 0.65 | 3264280 | 0.70 |
| 1:20 | 14295478 | 3.36 | 5081242 | 1.08 | 4575023 | 0.98 |
| 1:30 | 19389062 | 4.56 | 6969623 | 1.48 | 7331157 | 1.58 |
| 1:40 | 19709062 | 4.63 | 6590753 | 1.40 | 6759532 | 1.45 |
| 1:50 | 20313594 | 4.78 | 7161484 | 1.52 | 7257479 | 1.56 |

C3. Extraction method

| Solvent extraction | CUR |  | DMC |  | BDMC |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Area | Content | Area | Content | Area | Content |
| Sonication | 19949228 | 4.69 | 6668351 | 1.42 | 7078953 | 1.52 |
| Heat reflux | 17171444 | 4.04 | 5728892 | 1.22 | 6131291 | 1.32 |

C4. Number of extractions

| Number of extractions | CUR |  | DMC |  | BDMC |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Area | Content | Area | Content | Area | Content |
| Total extract 3 times | 16579980 | 3.90 | 5491998 | 1.17 | 5750372 | 1.24 |
| 4th extraction | 284420 | 0.07 | 86746 | 0.02 | 54909 | 0.01 |

| Solvent extraction A | Extraction method |
| --- | --- |
| Ratio of medicinal herb/solvent C | Number of extractions D |

Curcuminoid content was obtained during the optimization of extraction conditions

- METHOD VALIDATION

D1. System suitability

| 1 |  |
| --- | --- |
| 2 |  |
| 3 |  |
| 4 |  |
| 5 |  |
| 6 |  |

Data summary table of system suitability

| Analyte | BDMC | DMC | CUR |
| --- | --- | --- | --- |
| Retention time (min) |  |  |  |
| Mean | 9.353 | 10.343 | 11.448 |
| RSD% | 0.58% | 0.47% | 0.39% |
| Peak area |  |  |  |
| Mean | 3650550 | 4125305 | 13815704 |
| RSD% | 1.96% | 1.33% | 1.46% |
| Capacity factor |  |  |  |
| Mean | 8.353 | 9.343 | 10.448 |
| RSD% | 0.65% | 0.52% | 0.42% |
| Resolution |  |  |  |
| Mean | - | 2.805 | 2.823 |
| RSD% | - | 1.08% | 0.76% |
| Asymmetry |  |  |  |
| Mean | 1.167 | 1.108 | 1.117 |
| RSD% | 0.87% | 1.41% | 0.73% |
| Theoretical plate number |  |  |  |
| Mean | 12356.68 | 12469.00 | 12344.17 |
| RSD% | 1.24% | 1.37% | 1.80% |

a

b

c

e

d

D2. Specificity

(a) Mobile phase solvent; (b) Sample solvent; (c) Standard; (d) Sample; (e) Spiking

| Purity of curcumin | Purity of desmethoxycurcumin |
| --- | --- |
| Purity of bisdemethoxycurcumin | The 3D data graph of three analysis |

D3. Linearity of calibration curves

Data summary table of linearity of calibration curves

| No. | Concentration (µg/ mL) |  |  | Peak area (mAU.s) |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | CUR | DMC | BDMC | CUR | DMC | BDMC |
| 1 | 7.5 | 2.5 | 2.5 | 4651135 | 1674805 | 1644824 |
| 2 | 15 | 5 | 5 | 9462890 | 3466207 | 3367172 |
| 3 | 18 | 6 | 6 | 12071992 | 4387063 | 3901134 |
| 4 | 30 | 10 | 10 | 19552301 | 6759959 | 7254929 |
| 5 | 37.5 | 12.5 | 12.5 | 23755665 | 9221984 | 8660793 |
| 6 | 45 | 15 | 15 | 30026480 | 10923181 | 10796626 |
| 7 | 60 | 20 | 20 | 36035316 | 13317227 | 12936920 |
| 8 | 90 | 30 | 30 | 56215093 | 20575116 | 20375649 |

D4. Limit of detection (LOD) – Limit of quantitation (LOQ)

| LOD (n=3) | LOQ (n=6) |
| --- | --- |

| Limit of detection (LOD) – Limit of quantitation (LOQ) |  |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | CUR | DMC | BDMC |
| LOD (n = 6) | Conc (µg/mL) | 0.25 | 0.25 | 0.25 |
|  | S/N | 4.026 | 3.553 | 3.538 |
| LOQ (n = 6) | Conc (µg/mL) | 0.75 | 0.75 | 0.75 |
|  | S/N | 13.018 | 13.983 | 15.112 |

D5. Accuracy and precision

| Analyte | True Conc. (µg/mL) | Day 1 |  | Day 2 |  | Day 3 |  | Inter-day (n=9) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Area | Measured Conc (µg/mL) | Area | Measured Conc (µg/mL) | Area | Measured Conc (µg/mL) | Measured Conc (Mean ± SD, µg/mL) | RSD (%) | Accuracy (%) |
| BDMC | 6.4 (80%) | 4,692,307 | 6.96 | 4,185,655 | 6.20 | 4,183,628 | 6.20 | 6.46 ± 0.30 | 4.58 | 100.94 |
|  |  | 4,544,094 | 6.74 | 4,508,249 | 6.68 | 4,119,419 | 6.11 |  |  |  |
|  |  | 4,418,033 | 6.55 | 4,190,015 | 6.21 | 4,380,488 | 6.49 |  |  |  |
|  | 8.0 (100%) | 5,380,939 | 7.98 | 4,732,852 | 7.02 | 5,289,531 | 7.84 | 7.56 ± 0.38 | 4.98 | 94.53 |
|  |  | 5,454,572 | 8.09 | 4,908,191 | 7.28 | 4,944,209 | 7.33 |  |  |  |
|  |  | 5,055,368 | 7.49 | 4,877,427 | 7.23 | 5,267,900 | 7.81 |  |  |  |
|  | 9.6 (120%) | 6,103,091 | 9.05 | 6,113,813 | 9.06 | 6,120,755 | 9.07 | 9.22 ± 0.42 | 4.54 | 96.05 |
|  |  | 6,369,282 | 9.44 | 5,938,207 | 8.80 | 6,534,923 | 9.69 |  |  |  |
|  |  | 6,514,036 | 9.66 | 5,753,568 | 8.53 | 6,533,614 | 9.69 |  |  |  |
| DMC | 6.4 (80%) | 3,914,121 | 5.75 | 3,516,698 | 5.17 | 3,986,827 | 5.86 | 5.61± 0.25 | 4.54 | 87.68 |
|  |  | 3,806,283 | 5.59 | 3,878,689 | 5.70 | 4,018,159 | 5.90 |  |  |  |
|  |  | 3,590,395 | 5.27 | 3,745,210 | 5.50 | 3,923,719 | 5.76 |  |  |  |
|  | 8.0 (100%) | 5,259,965 | 7.73 | 5,189,389 | 7.77 | 5,728,609 | 8.42 | 7.87 ± 0.26 | 3.34 | 98.40 |
|  |  | 5,084,559 | 7.47 | 5,379,544 | 7.90 | 5,504,393 | 8.09 |  |  |  |
|  |  | 5,270,450 | 7.74 | 5,362,032 | 7.88 | 5,348,599 | 7.86 |  |  |  |
|  | 9.6 (120%) | 5,475,227 | 8.04 | 5,344,917 | 7.85 | 5,815,046 | 8.54 | 8.38 ± 0.38 | 4.52 | 87.32 |
|  |  | 6,014,505 | 8.84 | 5,809,782 | 8.53 | 5,886,640 | 8.65 |  |  |  |
|  |  | 5,610,619 | 8.24 | 5,393,033 | 7.92 | 6,007,336 | 8.83 |  |  |  |
| CUR | 24 (80%) | 15,664,559 | 25.41 | 15,008,195 | 24.35 | 15,759,285 | 25.57 | 25.06 ± 0.71 | 2.83 | 104.44 |
|  |  | 15,690,189 | 25.45 | 15,866,507 | 25.74 | 14,589,583 | 23.67 |  |  |  |
|  |  | 15,727,081 | 25.51 | 15,631,712 | 25.36 | 15,112,356 | 24.52 |  |  |  |
|  | 30 (100%) | 19,096,350 | 30.98 | 17,430,308 | 28.28 | 18,502,528 | 30.02 | 29.81 ± 1.41 | 4.73 | 99.36 |
|  |  | 19,563,594 | 31.74 | 16,982,585 | 27.55 | 18,968,842 | 30.77 |  |  |  |
|  |  | 18,948,710 | 30.74 | 17,631,683 | 28.60 | 18,235,341 | 29.58 |  |  |  |
|  | 36 (120%) | 22,781,069 | 36.96 | 23,734,494 | 38.50 | 22,875,795 | 37.11 | 37.59 ± 0.79 | 2.10 | 104.43 |
|  |  | 23,704,011 | 38.46 | 23,181,971 | 37.61 | 23,603,405 | 38.29 |  |  |  |
|  |  | 23,401,635 | 37.96 | 22,984,584 | 37.29 | 22,286,910 | 36.16 |  |  |  |

E. FINGERPRINT ANALYSIS

| No. | Sample code | Retention time |  |  | Content (%) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | CUR | DMC | BDMC | CUR | DMC | BDMC |
| 1 | CLR-01 | 11.3 | 10.2 | 9.2 | 2.05 | 1.21 | 0.11 |
| 2 | CLR-02 | 12.4 | 11.2 | 10.2 | 0.78 | 0.36 | 0.03 |
| 3 | CLR-03 | 12.3 | 11.2 | 10.2 | 1.13 | 0.61 | 0.09 |
| 4 | CLR-04 | 11.5 | 10.5 | 9.5 | 0.79 | 0.34 | 0.03 |
| 5 | CLR-05 | 11.8 | 10.7 | 9.8 | 5.68 | 2.71 | 2.52 |
| 6 | CLR-06 | 11.5 | 10.4 | 9.5 | 1.58 | 1.07 | 0.12 |
| 7 | CLR-07 | 11.2 | 10.1 | 9.1 | 2.14 | 6.92 | 1.29 |
| 8 | CLR-08 | 11.4 | 10.4 | 9.4 | 10.30 | 3.06 | 3.23 |
| 9 | CLR-09 | 11.5 | 10.4 | 9.5 | 0.96 | 0.56 | 0.07 |
| 10 | CLR-10 | 12.0 | 10.9 | 9.8 | 4.16 | 2.83 | 2.58 |
| 11 | CLR-11 | 12.0 | 10.9 | 9.9 | 9.89 | 2.69 | 2.47 |
| 12 | CLR-12 | 11.7 | 10.7 | 9.7 | 1.38 | 0.64 | 0.13 |
| 13 | CLR-13 | 11.8 | 10.8 | 9.8 | 0.77 | 0.33 | 0.03 |
| 14 | CLR-14 | 12.0 | 10.9 | 10.0 | 3.01 | 1.22 | 1.13 |
| 15 | CLR-15 | 11.6 | 10.5 | 9.5 | 5.06 | 2.31 | 1.90 |
| 16 | CLR-16 | 11.5 | 10.4 | 9.4 | 3.58 | 1.56 | 1.31 |
| 17 | CLR-17 | 12.4 | 11.2 | 10.2 | 10.16 | 2.60 | 2.57 |
| 18 | CLR-18 | 11.1 | 10.1 | 9.1 | 6.73 | 2.04 | 1.90 |
| 19 | CLR-19 | 11.2 | 10.2 | 9.2 | 5.46 | 1.90 | 1.70 |

CLR: Curcuma Longa Rhizoma