# Adapted NucleoMag Tissue Kit

## eDNA extraction protocol (400 µL)

Lysis

- Start the thermomixer and set it to 55 °C (program 1) or use the aluminium bead bath.

- Place the filter in a Falcon tube rack with the inlet facing upwards. Remove the inlet cap and add 100 µl of ProtK (optional: add up to 200 µL TNES if there might not be sufficient buffer) to each filter. Close the inlet cap and spin the filter 10 times to mix the sample.

- Incubate samples for 1h at 55 °C.

Chemicals

- Check if sufficient amounts of 10 ml syringes are available in the eDNA lab.

- Check if sufficient amounts of NucleoMag buffer aliquots are available in the eDNA lab:

- NucleoMag B-Beads (8 µl per sample)

- Buffer MB2 (640 µl per sample)

- Buffer MB3 (200 µl per sample)

- Buffer MB4 (200 µl per sample)

- Buffer MB5 (300 µl per sample)

- Buffer MB6 (55 µl per sample)

------------------------------------------------------------------------------------------------------------------------

Sample transfer

- Label one Eppendorf 1.5 ml tube for each filter.

- Remove both caps of the filter and place it sideways on the Falcon tube rack.

- Take a new 10 ml sterile syringe and the respective filter. Then, gently pull the sample out of the filter. Face the inlet downwards and the syringe upwards. Pull the syringe and wait until the sample stopped running into the syringe. Now gently remove the syringe without spilling droplets from the inlet.
Note: The sample volume can vary between 400 µl to 1200 µl depending on how much water was left before the buffer injection.

- Transfer the sample to the respective Eppendorf 1.5 ml tube by gently clearing the syringe. Prevent the occasionally produced foam from exiting the tube.

Clear lysate

- Set the thermomixer to 20 °C (program 2)

- Centrifuge the samples for 5 min at full speed (5,600-6,000 x g).

Bind DNA to NucleoMag B-Beads

- Add 8 µl of NucleoMag B-Beads and 320 µl Buffer MB2 to a new Eppendorf 1.5 ml tube.

- Add 200 µl of the cleared lysate and mix by pipetting up and down 6 times.

- Incubate for 5 min at 1000 RPM at 20 °C.

- Separate the magnetic beads for 2 min.

- Remove and discard the supernatant by pipetting (set pipet to 550 µl).

- Remove the tubes from the magnetic separator rack.

- Add 200 µl of sample and 320 µl Buffer MB2 to increase the extraction input volume to 400 µl and repeat steps 3-5 once.

Wash with MB3

- Remove the tubes from the magnetic separator rack.

- Add 200 µl Buffer MB3 and mix by pipetting up and down 6 times.

- Incubate for 5 min at 1000 RPM at 20 °C.

- Separate the magnetic beads for 2 min.

- Remove and discard the supernatant by pipetting (set pipet to 220 µl).

Wash with MB4

- Remove the tubes from the magnetic separator rack.

- Add 200 µl Buffer MB4 to each tube and mix by pipetting up and down 6 times.

- Shake the tubes for 5 min at 1000 RPM at 20 °C.

- Separate the magnetic beads for 2 min.

- Remove and discard the supernatant by pipetting (set pipet to 220 µl).

- Do not remove the tubes from the magnetic separator rack!

- Set the thermomixer to 56 °C (program 3).

Wash with MB5

- Leave the tubes in the magnetic separator rack!

- Gently add 300 µl Buffer MB5 to each tube without disturbing the pellet and incubate for 45 – 60 sec. Do not resuspend the magnetic beads in the Buffer MB5!

- Aspirate and discard the supernatant by pipetting (set pipet to 320 µl).

Elution

- Remove the tubes from the magnetic separator rack.

- Add 55 µl Buffer MB6 and mix by pipetting up and down 10 times.

- Incubate for 10 min at 1000 RPM at 56 °C.

- Separate the magnetic beads for 2 min.

- Transfer 50 µL of supernatant containing the purified DNA to a new tube.

- Store the samples at -20 °C until further processing.